## Slide 1: Instructor’s Electronic Curriculum Resource For Techniques in Noninvasive Vascular Diagnosis-4th edition. by Robert J. Daigle, BA, RVT, FSVU Produced by Summer Publishing. LLC
This collection is designed to aid instructors and students that are using the textbook Techniques in Noninvasive Vascular Diagnosis. Summer Publishing authorizes and licenses the use of this COPYRIGHTED material contained in this collection to the school educational program. Use of the slides, graphics, case studies or movies for endeavors outside the activities of the school is prohibited without written permission from Summer Publishing.

## Slide 2: Technique in Noninvasive Vascular Diagnosis-4th Ed. An Encyclopedia of Vascular Testing
Textbooks be purchased at www.summerpublishing.com
Info@summerpublishing.com

## Slide 3: Instructor’s Electronic Curriculum Resource- Hyperlinks have been incorporated into images throughout the slide collection. You must have an internet connection for hyperlinks to work. It’s also useful to have a browser open and running on you PC while in slide show mode. Alternatively, the movies can be re-inserted into PowerPoint . After viewing the You Tube video demo, return to Powerpoint, select Presenters View or Slideshow to continue.

## Slide 4: Chapter 1. Principles of Doppler Ultrasound Techniques In Noninvasive Vascular Diagnosis - 4th edition
By Robert J Daigle, BA, RVT, FSVU

## Slide 5: Definitions for this chapter
Frame rate-
  “real-time” ultrasound is generated by a series of still frames that are displayed in rapid succession, creating motion. Frame rate is expressed as the number of frames per second (FPS). Motion pictures use upwards of 24 fps
Notes: Definitions are also in the Glossary in the book. But students may not know some of the basic terminology.

## Slide 6: Definitions for this chapter
Temporal resolution-
  the ability to detect that an object has moved over time. Related to frame rate.
Spatial resolution-
  defined as the ability to distinguish two points as separate in space in any direction
Artifact-
  a substance or structure not naturally present in the matter being observed but formed by artificial means.
Notes: Definitions are also in the Glossary in the book. But students may not know some of the basic terminology.

## Slide 7: Definitions for this chapter
Spectral Broadening-
  a filling-in and widening of the spectral waveform
Range gate-
  same as sample volume (sv)
Bi-directional-
  the ability to display forward and reverse flow directions.
Notes: Definitions are also in the Glossary in the book. But students may not know some of the basic terminology.

## Slide 8: Definitions for this chapter
Cosine theta-
  in Doppler ultrasound, the numeric value of the angle (theta) at which the ultrasound beam intersects with the axis of flow. E.g., the cosine of 60º is 0.5
B-mode ultrasound-
  brightness mode ultrasound, aka, 2-D in cardiac ultrasound nomenclature.
  the image is created by pixels of different shades of grey or “brightness”, based on the strength (amplitude) of the reflected energy.
Notes: Definitions are also in the Glossary in the book. But students may not know some of the basic terminology.

## Slide 9: Definitions for this chapter
Scanline-
  a line or path of the ultrasound beam, and subsequent echo. Multiple scanlines, generated by multiple piezo elements, are used to create an image “frame”.
Notes: Definitions are also in the Glossary in the book. But students may not know some of the basic terminology.

## Slide 10: Doppler Shift
Light or sound from moving objects will appear to have different wavelengths depending on the relative motion of the source and the observer.
Notes: In 1842, Christian Andreas Doppler presented his hypothesis “On the coloured light of the double stars and certain other stars of the heavens”. He postulated that certain stars appeared red or blue, when seen through a telescope, because they were in motion away or towards the observer on earth. Stars emit light waves and the frequency of the light waves is changed due to the motion. If the star is moving towards the earth the lightwaves are of a shorter wavelength (thus a higher frequency) (more cycles per second) (blue-shifted). On the visible light spectrum, blue has a higher frequency than red.
Conversely, if the star is moving away from the earth, the light waves reaching the observer on earth are of a lower frequency (red-shifted). In the visible light spectrum, red has a lower frequency than blue. These same principles apply to sound waves.
The color Doppler color coding of “Blue away-Red towards contradicts actual physics. Color Doppler was developed for cardiac applications before peripheral vascular; the predominate cardiac flow was “towards” the transducer, so it was arbitrarily assigned red color, with retrograde flow (regurgitation) (away from transducer) being assigned blue.

## Slide 11: Doppler Shift
If a sound source moves rapidly towards the observer, audible frequency (pitch) will increase
If a sound source moves away rapidly, frequency, or pitch, will decrease
Same goes for light waves.

## Slide 12
Formula 1, race start
No movement - no Doppler shift
Sound waves
Notes: From this Formula One race car, sound is emitted in all directions, and frequency is similar in all directions if the car is stationary.

## Slide 13: Doppler Shift
120 MPH
200 MPH
Frequency increases in front of car and decreases as it moves away.
The faster the car goes, the greater the frequency “shift”.
Notes: When the car is speeding the emitted sound waves are compressed in front of the car, and elongated as it moves away. The pitch of the sound increases as the car approaches you, and the pitch decreases as the car moves away. You are hearing the Doppler Frequency Shift!

## Slide 14: Doppler in Vascular Diagnosis
Blood cells don’t emit sound.
So, we must transmit a sound and evaluate the echo.
The difference between the transmitted frequency and the reflected frequency (the echo) is called the Doppler shift
Imaging and Doppler ultrasound use Piezo-electric effect

## Slide 15: Piezoelectric elements vibrate when voltage is applied:
Piezo-electric
“piston” effect
Sound waves
Electrical energy Mechanical energy
voltage
Conversion of electrical energy into mechanical.
Notes: When elements (groups of elements) in the ultrasound transducer are stimulated with electricity, they expand and contract (oscillate) and produce a high frequency sound signal. As in a sound speaker, it’s a conversion of electrical energy into mechanical energy , (the moving speaker baffle).

## Slide 16: Returning sound waves vibrates the piezo elements
Piezo-electric
“piston”
Echo sound wave
Voltage
Mechanical energy Electrical energy
Conversion of mechanical energy into electrical.
Notes: Returning ultrasound echoes (sound waves) mechanically stimulate and move the crystals: the mechanical motion is converted into electrical current and stored. A similar scenario occurs with your tympanic membrane in the ear; it is mechanically oscillated by sound waves, and the energy is converted into an electrochemical signal and sent to the brain.

## Slide 17: Frequency
The frequency of the transmitted signal depends on the size (thickness) of the elements: the smaller the elements, the higher the frequency.
Frequency is expressed in hertz, 1 sound wave or cycle per second = 1 hertz.
For diagnostic ultrasound, Doppler frequency is between 2.0 - 10 MHz

## Slide 18: Bi-directional Doppler
Quadrature detection and processing allows “directionality” to be determined.
A simplified description: the transmitted frequency is subtracted from the returning frequency.
If the returning frequency is higher, the Doppler shift is “positive” (+) shift
If the returning frequency is lower, the Doppler shift is “negative” (–) shift.
Notes: In Bidirectional Doppler, positive shift is flow TOWARDS the Doppler Beam, and in the spectral or analog display, shown above baseline. Flow AWAY from the Doppler beam is negative flow, and displayed below baseline.

## Slide 19
transmit
return
transmit
return
F = Fr - Ft
piezoelectric crystal
The Frequency Shift
∆F (delta F) = frequency shift
Fr = returning frequency
Ft = transmitted (incident frequency)
Notes: Doppler shift is the difference between the returning and transmitted frequencies.
Flow towards Doppler beam increases return the frequency, Flow away decreases frequency (fewer cycles per second).
Flow away from Doppler beam is a “negative” shift and is displayed below baseline.
Frequency shift values will be similar for flow towards and flow away if the velocity and Doppler angle are the same.
NOTE: transmitted (incident) frequency is subtracted from the returning frequency-
If transmitted frequency (Ft) is 5.0 MHz, and the returning frequency (Fr) is 4.95 MHz the frequency shift is 4.995 – 5.0 = – 0.005 MHz or -5.0 KHz (negative 5 KHz), so flow is away from the Doppler beam. (if you subtract a larger number from a smaller number, you get a negative or minus value).
Conversely, if the returning frequency (Fr) is 5.005 MHz, the frequency shift is 5.005 – 5.0 = 0.005 MHz or 5.0 KHz (“positive” 5 KHz), so flow is towards the Doppler beam.

## Slide 20
transmit
return
transmit
return
F = Fr - Ft
piezoelectric crystal
The Frequency Shift
If the speed of the blood is the same in figure A and B, the frequency shifts will be the same, but one will be “negative, and the other “positive” to indicate flow direction
A
B
Notes: Doppler shift is the difference between the returning and transmitted frequencies.
Flow towards Doppler beam increases return the frequency, Flow away decreases frequency (fewer cycles per second).
Flow away from Doppler beam is a “negative” shift and is displayed below baseline.
Frequency shift values will be similar for flow towards and flow away if the velocity and Doppler angle are the same.
NOTE: transmitted (incident) frequency is subtracted from the returning frequency-
If transmitted frequency (Ft) is 5.0 MHz, and the returning frequency (Fr) is 4.95 MHz the frequency shift is 4.995 – 5.0 = – 0.005 MHz or -5.0 KHz (negative 5 KHz), so flow is away from the Doppler beam. (if you subtract a larger number from a smaller number, you get a negative or minus value).
Conversely, if the returning frequency (Fr) is 5.005 MHz, the frequency shift is 5.005 – 5.0 = 0.005 MHz or 5.0 KHz (“positive” 5 KHz), so flow is towards the Doppler beam.

## Slide 21: Flow Away from Doppler Beam
CCA Flow
Invert
CCA Flow
Manual Invert
Notes: The waveform, the “negative or minus” signs, and the color are shifted during INVERT.

## Slide 22: Auto - invert, spectral and color
Negative
Positive
99.3 cm/s
Left steer
Right steer
Positive
Negative
- 94.2 cm/s
Notes: Many ultrasound systems employ “auto-invert”; when the color box is steered in the opposite steering direction, the color coding does not change. However, the color bar color switched. Ditto for spectral Doppler, the waveform remains on the same side of the baseline, but the negative and positive signs on the scale bar flips.

## Slide 23: HP-Philips “Intelligent Doppler”
Dial arrow cursor to point in the direction of flow; antegrade flow
will be above baseline, retrograde flow below baseline, regardless
of steering direction.
Notes: Hewlett Packard invented a unique method of direction determination. This technology is now incorporated into a number of Philips Ultrasound systems (they acquired HP).
The arrow on the angle adjust cursor should be pointed in the direction that flow is supposed to be going. This applies for arteries as well as veins. Antegrade flow will always be above baseline, regardless of steering direction. Retrograde flow will be below baseline. Don’t use the invert function, or it will defeat this simple concept.

## Slide 24: Frequency Shift is Related to:
Transmitted Doppler frequency
Velocity of blood cells
Angle of Doppler beam to flow
  0 º = highest F shift, 90º = lowest F shift (worst Doppler angle)
Notes: The faster the moving red blood cells, the higher the frequency shift. The higher the transmitted Doppler freq, the higher the frequency shift.

## Slide 25: Doppler Angle versus Freq. Shift
45
60
90
60
45
flow towards Doppler beam
flow away from Doppler beam
linear array transducer
Frequency shift increases with smaller (numeric) angles
Notes: The smaller the Doppler angle to flow, the higher the frequency shift.

## Slide 26
Frequency Increase Due to Angle
linear array transducer
30
60
INVERTED
A
B
Assume velocity at point
A + B is the same
Notes: Assume no change in the velocity within this vessel, and no change in steering angle. Point “A” will have a higher frequency than “B”, because the angle to flow is smaller (30º versus 60º)..

## Slide 27
linear array transducer
A
B
Aliasing at region “A” because the frequency is higher that at “B”. The Doppler angle at A is smaller.
Green lines are drawn in parallel to the Color Doppler scanlines to
demonstrate the angle changes at A versus B.
Notes: Higher frequency at “A” compared to “B” due to lower angle to flow at “A”.

## Slide 28: DOPPLER EQUATION
f =
C
2 ft V Cos
∆f = frequency shift
2 x ft = 2 x transmitted frequency
V = velocity
Cos Θ = Cosine theta-Doppler angle of incidence
C = average speed of sound in tissue
Notes: Blue figures are factors that are programmed into the ultrasound microprocessor, ie, transmitted Doppler frequency and average speed of sound in tissue (1540 M/sec).

## Slide 29: Doppler equation to calculate velocity
V =
f
C
2 ft Cos
Δf = frequency shift
2 x ft = 2 x transmitted frequency
V = velocity
Cos Θ = Cosine theta-Doppler angle of incidence
C = average speed of sound in tissue
Blue values are programmed into the ultrasound system
Notes: Formula rearranged to solve for Velocity.

## Slide 30: Mnemonic tip
F =
C
= 3 other things
V
2 ft
Cos
Notes: Screwy, but if you think of calf spelled as CAF, or C∆F, that’s half of the items in the equation.

## Slide 31
V
=
F
Cos
KHz
transducer
Notes: Simplified Doppler equation: User must provide the angle to flow (Cos ) and the frequency shift to derive a velocity estimate.
From the equation, if angle changes but frequency is constant, velocity value will change.
If angle is constant, but frequency changes, velocity change will have occurred.

## Slide 32: Frequency is the same in A, B, C. Notice how the Velocities are different due to different cosine theta angles
V
=
F
Cos
Notes: In figures B and C below, the angle cursor was arbitrarily adjusted (misaligned) to other angles on the same image, 34 and 79º respectively. Notice that the frequency does not change , it's 2.83 kHz in each instance, only the velocity calculation changes (65.8 cm/s and 268 cm/s). In the equation above, when the recorded frequency remains constant, if you change the cosine value (the Doppler angle), the velocity calculation must change.

## Slide 33: Cardiac Doppler is a vertical world
Flow is directly towards or away from Doppler beam- SIMPLE !
Towards
Away

## Slide 34: Vascular is a “horizontal world”

## Slide 35: Doppler Electronic Steering: Doppler beam is “phased” to provide a non-perpendicular angle
Notes: If the Doppler beam is not phased, it is “fired” straight down and parallel to the sides of the ultrasound image. There is “Zero” degrees of beam steering. Fortunately, manufacturers allow the beam to be steered to the left or to the right by “phasing”. The phasing is called steering. If the beam is unsteered, the beam steering angle is 0º. The maximum amount of steering is about 30º. This is advantageous to acquire a good Doppler angle to flow. When manufacturers try to phase the beam beyond 30º, a “grating lobe” artifact appears in the spectral display.
Low-end, economical ultrasound systems can be steered to only 15º.
High frequency transducers usually can only be steered to 20º.
This will require “heel and toe” transducer manipulation to obtain a 60º Doppler angle.

## Slide 36: Steering Angle versus Angle of Incidence
Notes: Maximum steering angle on most ultrasound systems is 30º. High frequency transducers can only steer to 20º. Assume a vessel is parallel to the skin, as in this image, the Doppler angle will be 90 degrees if there is no steering (zero degrees of steering). If the beam is phased (steered) to 30 degrees, the angle to flow will be 60 degrees. A 60 degrees angle to flow, (The Doppler angle), or less, is for velocity calculation.

## Slide 37
Ice fisherman with his fish line dropped straight down through the hole (zero degrees of line angle). The bottom angle is 90 degrees.
A turtle grabs the bait and walks a few feet along the bottom. The line angle has changed by 20º from where it was. The bottom angle is now 70 degrees.
Turtle pulls the bait several more feet along the bottom. The line angle is now 30º from where it was originally at zero. The bottom angle is now 60 degrees.
Steering Angle Explained
Notes: A cartoon to help with the confusion of Beam steering angle, versus Doppler angle to flow. The fishing line angle is steering angle, the bottom angle is Doppler angle to flow.

## Slide 38: Doppler Beam Steering
Purpose: to achieve a 60º angle to flow (or less).
  AKA “angle Theta”
  AKA-the Doppler angle

## Slide 39: Doppler Flow Direction & Steering
Popliteal art.
ICA
Flow
Popliteal artery
Notes: Beam steering is altered frequently during most vascular exams.

## Slide 40
steered 20 o
linear transducer
60 o
To obtain Doppler angles of 60º or less, Steer to the side that’s deeper.
steered 20 o
linear transducer
80 o
Bad !
steered 20 o
linear transducer
60 o
Good!
Good!
steered 20 o
linear transducer
80 o
Bad !

## Slide 41: Doppler angle cursor must be aligned to the axis of flow

## Slide 42: “Heel & toe” transducer adjustment to decrease Doppler angle to flow
Notes: If a transducer only steers the Doppler beam to 20º, the operator must create a lower angle to flow by manipulating the transducer in a “Heel & Toe” method, or steering in the most advantageous direction.

## Slide 43: “Heel & Toe” to Improve Doppler Angle
Notes: If the Doppler beam can only be steered to 20 degrees (figure A and B), then heel & toe maneuver (figure B) must be used to achieve a 60 degree angle to flow.

## Slide 44: Popliteal Artery Steering Demo
AC 60
Bad steering- 90º
Good steering

## Slide 45: Both “Read” 60 Degrees!
AC 60
AC 60
Bad steering and bad
angle alignment!
Notes: Although the angle read-out indicates 60 degrees, this is NOT a 60 degree angle to flow! The angle cursor is misaligned.

## Slide 46: Angle Cursor Misalignment
Actual Velocity = 100 cm/s
5 degree inadvertent misalignment
5º off at 40o = vel. calc = 108 cm/s
5º off at 60o = vel. calc = 118 cm/s
5º off at 80o = vel. calc = 195 cm/s !!
Perfect alignment,
Velocity = 100 cm/s
Notes: Demonstrates the theoretical velocity calculation errors of a 5 degree angle correction misalignment at various Doppler angles.
At a 40 degree angle, a 5 degree misalignment causes an 8% variation or error. At an 80 degree angle, a 5 degree misalignment causes the velocities to nearly double. There are other factors that cause velocity over estimation and these will be discussed later. Doppler angles of 60 degrees or greater can be used, BUT velocity calculations will be overestimated and in error.

## Slide 47
+
Doppler angle
Frequency
shift
VELOCITY
Notes: Correct!

## Slide 48
No angle cursor, system
assumes it’s zero degrees.
Velocity measurements will
be WRONG
Angle cursor aligned to axis
of flow. Velocity estimates
are OK

## Slide 49: Flow Velocity Estimation Purpose:
Quantitation of arterial stenosis
Define pathologic conditions
Typical measurements include peak systolic velocity (PSV) and end-diastolic velocity (EDV)
Frequency shift and Doppler angle to flow are essential ingredients

## Slide 50: Curved Array Doppler
The Doppler beam is unsteered in curved array technology.
It’s angled in the image due to the curve of the array.
Notes: The red line drawn in at the top of the Doppler line is perpendicular to the Doppler line. This means zero (no) beam steering. The relationship will be maintained when the Doppler line is moved anywhere along the curve of the transducer. The angle is created by the curvature of the transducer, not phasing.

## Slide 51: Doppler Devices
Continuous-wave (CW) Doppler
  2 piezo-electric crystals: one transmitting, one receiving
  No aliasing artifact
  Not specific for depth
Notes: The frequency of the transmitted signal depends on the size of the elements: the small the elements, the higher the frequency.

## Slide 52: Analog Doppler
Common form of CW Doppler
Zero-crossing detection method
Displays instantaneous average frequency shift
Not depth specific,
No sample volume

## Slide 53: Analog Doppler
If venous and arterial signals are present in the beam path, they’re averaged together, and waveform distortion can occur
“Distortion” of analog waveform due to
arterial and venous signal averaging.

## Slide 54: Doppler Devices
Pulsed-waved (PW) Doppler
  single element: transmitting, then receiving (used for TCD)
  multi-element linear arrays
Piezo elements are
“diced” on ceramic
Tiny wires to each element
Curved array

## Slide 55: Pulsed-Wave Doppler
Alternating pulse transmission and reception
Transducer design
  single element transducers (TCD)
  multi-element linear array
  phased array transducers
Sample volume (range gate, sample gate) are integral to pulse Doppler systems

## Slide 56: Sample Volume
Sample volume length determines the transmit pulse length.
Pulse length is the “ring time” or duration of transmission.
One “pulse” transmission with 4 cycles
SV

## Slide 57: Sample Volume
Transmitted Doppler pulse contains numerous cycles (5-30 Hz)
The number of cycles in a pulse multiplied by the wavelength = spatial pulse length
Large SV,
long pulse length
Small SV,
short pulse length

## Slide 58: RANGE-GATED PULSED DOPPLER
1540 M/s = 1 cm /6.7 microseconds
If sample volume at 4 cm, pulse cycle = 54 µ sec
Notes: 1.The sample volume is set to a specific location and a pulse (burst) is transmitted from the piezo elements.
2. The transmit and receive Doppler processing circuits are turned off.
3. The pulse echo reflects back to the elements from various depths, and perhaps from intervening vessels, but the signals are ignored as the receive function is not "live".
4.The transmitted pulse reaches the sample volume depth and reflects back. Portions of the energy continue deeper into tissue.
5. At a specific time interval, THAT IS DETERMINED BY THE DEPTH OF THE SAMPLE GATE, the receive mode is activated for a brief period to receive an echo from that specific depth. The frequency component of the echo is stored and processed. The frequency content of the returning pulse is then compared to the transmitted signal for frequency shift.
6. The receive mode is turned off, and a new pulse is transmitted. The transmit - receive cycle is then repeated.

## Slide 59: Pulse Repetition Frequency (PRF)
Number of pulse-echo cycles per second
Expressed in hertz (Hz)
Aliasing occurs when Doppler shift frequencies exceed 1/2 of PRF
PRF is adjusted with “scale” control.
Notes: The number of times the pulse transmission, wait, receive cycle occurs per second is called the Pulse Repetition Frequency (PRF)

## Slide 60: Spectral Aliasing
3.0
8
Notes: On the left, the PRF is too low, aliasing occurs and wraps back up into the spectra. On the right the PRF has been increased to eliminate aliasing.

## Slide 61: Aliasing
Notes: Book graphic page 8.

## Slide 62: Aliasing Remedies
OR
PRF 4.34 KHz
Notes: To reduce aliasing, increase PRF (scale), or drop the spectral display baseline. A theoretical solution : Decrease transmitted Doppler frequency;

## Slide 63: Fast-Fourier Transform (FFT)
Samples the Doppler signal for frequency components
Displays:
  Doppler shift frequencies (vertical axis)
  Time (horizontal axis)
  Amplitude or power ( brightness of pixels)

## Slide 64
Doppler Guidelines
For accurate velocity measurements use angles of 50-60 degrees.
Obtain multiple samples from stenosis
For carotid follow-up exams
  use same Doppler angle in stenosis
  use same Doppler steering direction

## Slide 65: FFT Amplitude “bins” displayed in gray scale
“bins”
bright
Frequency
scale
Frequency shifts grouped into bins
dark
gray
dark
gray
black
black
Notes: Not all frequency shifts can be displayed in any one moment of time, so they are grouped into bins of similar frequency shifts. The higher the number of similar frequency shifts in a bin, the brighter the pixel on the FFT display. The peak (highest) frequency display and lowest frequency display will be a darker shade of gray, as there are fewer blood cells traveling at these speeds (Remember parabolic flow patterns.) The highest density of blood cells at a given speed, or rather the highest number of frequency shifts in a specific range will be assigned a bright shade of gray in the displayed pixels. This is called the Doppler amplitude or Doppler Power. A Color scale can be assigned to the amplitude instead of gray scale.

## Slide 66: Spectral Spread
Peak frequencies
Mean frequencies
Lowest frequencies
In any one moment of “Doppler” time a bizillion red blood cells pass through the SV.
There is a range of RBC velocities, they all don’t move at the same speed

## Slide 67: Spectral and Color Doppler Controls: Filters
Wall filters
Also called High pass filters
Also called Thump filters
Notes: All are names for wall filters. Wall filters eliminate low frequencies along the spectral baseline or color baseline.

## Slide 68: Wall filters are essential in eliminating low frequency noise, but if set improperly, can adversely affect waveform morphology.
Low filter, 55 Hz
High filter, 200 Hz

## Slide 69: COLOR DOPPLER: A color image frame is obtained by using multiple color scan lines
Notes: Multiple scan line are used to acquire color information.

## Slide 70: Color information is obtained by Autocorrelation
Initial pulse transmitted along scan line
Received echo is stored
Subsequent pulse - echoes are stored and compared to initial reference echo
3 pulses required, 8 - 20 commonly used
Notes: Autocorrelation method:
An initial pulse is transmitted down a single scanline and into the tissue. The non-Doppler-shifted information is processed for gray-scale content, while the Doppler-shifted echoes from the multiple sample sites along the scan line are stored in memory. The echoes are converted to voltages for each echo.
Subsequent pulse echoes are obtained, stored and then compared to the initial pulse echo.
Shifted echoes (voltages), related to motion within the sample locations, are detected by the comparison

## Slide 71: Autocorrelation Derives:
Mean frequency shift
Flow direction
Variance (spread of frequencies)
Power
Notes: Mean frequency shift, not peak frequency.

## Slide 72: Pulses per Scanline
B - Mode = 1
Color = 3 - 20 (commonly 8 - 10)
Packet size
Ensemble length
Color sensitivity
Color quality
aka.
Notes: Color autocorrelation requires a number of pulses to se sent along or down each scan line. Image information and color info is obtained in alternate fashion. The highr the number of pulses per line (packet size) the better the color, but at the expense of frame rate.

## Slide 73: Pulses per Scanline
High number = robust, full color,
but lower frame rate
Low number = weak, “holey” color,
but high frame rate

## Slide 74: Color Scanline Density
Small pixel size
“Brick” shaped pixels
High scanline density
Low scanline density
Good spatial resolution
Slow frame rates
Reduced spatial resolution
Faster frame rate
Notes: High scanline density provides good color spatial resolution but requires more time resulting in a slower frame rate.
Low scan line density compromises resolution but improves frame rate. Manufacturers often “fill-in” the spaces between scanlines; this is called spatial averaging.

## Slide 75: Spatial Averaging
Manufacturers use spatial averaging to improve frame rates.
In the image on Rt. the averaging “went south”!
Notes: Temporary malfunction in color processing discontinued spatial averaging, allowing the scanlines to be seen.

## Slide 76: Color Frame Rate
Pulses per line (packet size)
Density of scan lines
Color box width
PRF

## Slide 77
Triphasic flow
Notes: In lower extremities, color flow depicts multiphasic flow similar to that seen in spectral waveforms. Frame rates must be fast in order to capture the rapidly changing flow direction.

## Slide 78: Frame rate versus color box size
FR 29 Hz
FR 10 Hz
Notes: Slow frame rate, particularly in the lower extremity arteries, can cause a “strobe-light” effect, and can result is poor vessel filling.

## Slide 79: Artifact related to slow frame rate
Notes: The blue color in the image on the left is residual “lag time” from the reverse flow componetn of the popliteal artery. If Frame rate is sufficient (image on right), the entire artery segment within the color box should be red, then blue, then briefly red demonstrating triphasic flow. If the frame rate is slow, the blue reverse flow color remains on the screen while the red forward flow is demonstrated.

## Slide 80: Flow Direction Primer
Flow relative to Doppler beam (or laser radar) is either Towards, Away, or Perpendicular
Notes: Police RADAR uses the Doppler principle to measure velocity of a moving object

## Slide 81
Flow Direction Primer
Notes: The car represents red blood cells. Doppler shift is used to detect speed and direction of blood cells and cars. Doppler shift does not work at angles close to 90º, nor does police RADAR

## Slide 82
Color Coding Flow Direction
BART
Blue Away Red Towards
Color invert = Red Away
AWAY
TOWARDS
Notes: Standard color orientation is “BART”

## Slide 83
Notes: All flow Towards Doppler scanlines will be colored Red. All flow away Blue.

## Slide 84
Notes: All flow Towards Doppler scanlines will be colored Blue. All flow away Red.

## Slide 85
Notes: Flow at 90º to Doppler color scanline will be poorly coded, ambigously color coded.

## Slide 86
Notes: Note, the peroneal artery is blue, as it is “Flow Away:

## Slide 87: Color Angles Game- Mid Term Exam!
Is the gastroc vein flow normal (into the popliteal vein), or abnormal (indicating reflux)?
Color is inverted
ie, red is “away”.
Feet >>
<< Heart
Notes: Here’s a color game test: Is the gastrocnemius flow, that is supposed to be draining into the Popliteal vein, traveling in the correct direction, or is it reflux flow (coursing back into the gastroc vein)???
System is set for blue towards, red away, ie, color is inverted.
Gastroc vein flow direction is normal, it is RED, so it is AWAY from the Doppler beam. The popliteal artery is also red and away. Remember, color coding is based on flow relative to the Doppler beam, not whether it’s going left or right in the body. If the gastroc flow “color” was blue, if would be abnormal reflux flow.

## Slide 88
Color Steering
Popliteal artery
Unsteered
Steered right
Why is the color so lousy?
Notes: It’s fascinating the see how many people can perceive the problem on the left image, but not the right. It’s obvious that the angle to flow is 90 degrees on the left image, which is also the scenario on the right image. But many will be confused by the steering and not perceive the 90 degrees to flow.
First 2 images are at 90 degrees. One must use the side line on the color box to judge angle to flow. The top and bottom lines only define the limits of the color box, they can confuse the angles game, so ignore them.

## Slide 89
Color Steering
Popliteal artery
Notes: Color is steered correctly in this image.

## Slide 90: Learn to read the display and identify Doppler frequency
Notes: Learn to locate where the frequency indication is displayed on your screen.

## Slide 91: Color Scale (PRF)

## Slide 92: Color Gain
Notes: Tip: When adjusting color gain, turn it up until color blooms out of the vessel, then dial it back
so that it's confined within the walls; this way you know that you have adequate gain.

## Slide 93: Use aliasing to identify the “jet”

## Slide 94: Color Angles Game
Explain the blue colors.
Notes: Why are these various segments blue?

## Slide 95: Absence of color fill due to high “local” Doppler angle and PRF
PRF lowered
Notes: Is this non-filling area plaque?
The “local” angle on the left image is close to 90º, also the flow in this segment may be slow. The combination causes the frequency shift to be low, and flow not to be detected. In the image on the right , the PRF scale has been lowered and flow was revealed.

## Slide 96
Notes: PTA and Peroneal are both flowing in the correct direction.

## Slide 97: Aliasing vs. Reverse Coding
Nyquist limit
Notes: Aliasing is displayed as a change in color gradient. A change is color (red versus blue) that is separated by a black line usually indicates a change in direction relative to the transducer. See next slide.

## Slide 98
Aliasing vs. Reverse Coding
Notes: Note black line.

## Slide 99: Color Doppler Frequency
Locate the transmit frequency indicator on your system

## Slide 100: Doppler Penetration- 6 MHz
Notes: This 6 MHz frequency does not have adequate penetration for this vertebral artery

## Slide 101: Doppler Penetration- 4 MHz
Notes: 4MHz Doppler penetrates better than the 6 MHz Doppler = good filling.

## Slide 102: Caveat! Most color duplex systems lose sensitivity when steered.
Doppler steered
Doppler unsteered
Loss of sensitivity and penetration when beam is steered
Notes: Color and spectral Doppler can be used unsteered if the vessel dives downward. Most ultrasound Doppler systems lose sensitivity and penetration when steered to left or right. They are most sensitive when “fired” straight down. This is demonstrated in the images here. The loss of sensitivity when Doppler is steered only becomes apparent in deep vessels; and it varies by ultrasound system design.

## Slide 103: Power Doppler
Ultrasound Angio tm
Color Doppler Energy
Color Power Angio
Power Doppler Imaging
Convergence Power Doppler

## Slide 104: Power Doppler: advantages for flow in tissue
Color encoding of Doppler “amplitude” (density of similar velocity blood cells

## Slide 105: Power Dopper- organ perfusion

## Slide 106: Power Doppler: advantages for flow in tissue
color encoding at 90 o
more sensitive than color Doppler
good wall definition (non-moving walls)
no aliasing
many systems do not encode for flow direction

## Slide 107: Power Doppler- angle independent
Power Doppler
Color Doppler
Notes: Power Doppler works better than color Doppler at 90º to flow.

## Slide 108: Good Filling at 90º

## Slide 109: Effective Use of Power Doppler
Near total occlusion
Tortuous vessels- distal ICA
Notes: Power Doppler has been recommended to help identify low flow in “trickle” flow, near total occlusion, situations. Also, in tortuous vessels, you don’t get the multiple color changes of red, blue. red, over the course of the vessel.

## Slide 110: Power Doppler: Current Limitations
wall definition - “bleeding”
high persistence - slow frame rate
cannot detect increased velocity
differentiation of arterial /venous flow,
flow direction (some have directional ability)

## Slide 111: Disadvantage: Power Doppler removes hemodynamic information
Power Doppler
Color Doppler
Notes: In this combination of images, it is clear from color Doppler that the stenosis is causing a velocity increase and post stenotic turbulence, I.e., a hemodynamically significant stenosis. This is not evident in the power Doppler image.

## Slide 112: Power Doppler Clinical applications
Tissue or organ perfusion - yes
  native kidney, liver, transplants,
Vascularity in tumors - yes
Transcranial Doppler - yes
“Trickle flow” - yes
Arterial stenosis - no

## Slide 113: Artifacts

## Slide 114: Mirror Image Artifact of Proximal CCA
CCA
Artifact
Notes: Mirror images are often seen in the subclavian artery region. They can also be seen occasionally in the CCA. It is sometimes referred to as the “ghost” CCA.

## Slide 115: Mirror Image Artifact
A
B
C
hard reflective
surface
mirror image
object
C
D
Linear transducer
Linear transducer
object
A
B
1.
2.
Notes: 1
In figure 1. the object is displayed at a depth that corresponds to the “time of flight” of the transmitted and reflected ultrasound signal.
In Figure 2. the scanline “A” and reflection “B” accurately place the object within the scan field of view. However, other beams (C) are transmitted from adjacent elements along the array. Scanline “C” hits a reflective surface and the beam scatters. Some of the signal reflects off the object, back to the “hard” reflective surface and back to the array elements (D). Because the time of flight is longer, the object is displayed (or ghosted) at an increased depth. So two objects are displayed instead of one. The object and artifact are not always displayed in the same plane.

## Slide 116: Other Mirror Image Artifacts:
Left subclavian artery
Reflection artifact
Notes: The subclavian artifact is caused by reflection of the ultrasound beam off the air-tissue interface of the lung. The pedal artery artifact is due to reflection off the adjacent bone.

## Slide 117: Elevation Plane Focus Artifact

## Slide 118: Lateral Resolution
Lateral resolution.
The ability to distinguish two closely
spaced objects in the horizontal plane
Linear array transducers display a two
dimensional image of a 3-dimensional object.

## Slide 119: Dynamic Focusing
Transducer
Multiple focal zones
improve lateral
resolution
at multiple depths.
Click on image to start animation

## Slide 120: Dynamic lateral focusing provides multiple focal zones at depth

## Slide 121: Axial resolution is determined by spatial pulse length, which in turn is dictated by transducer frequency
Good axial resolution: the ability to distinguish adjacent structures in the vertical axis.

## Slide 122: The Third Dimension- the elevation plane
AKA – slice thickness
“slice thickness” varies with depth.
focal point is fixed and determined by transducer design and frequency

## Slide 123
If the slice thickness is wider than the object, the surrounding tissue may be displayed along with the object of interest.
vessel

## Slide 124: 5 MHz transducer with mid-elevation focal depth
GSV
Partial volume effect
Notes: No “filling” in transverse plane, but vessel appears filled in long view due wide slice thickness in that region.

## Slide 125: High frequency transducer with short (superficial) elevation focus
GSV
Notes: Again, no filling in the GSV in transverse. In long view there is no filling due to the vessel being in the focal zone in the elevation plane.

## Slide 126
“Matrix” transducers can provide a more uniform elevation focal zone.

## Slide 127: Intrinsic Spectral Broadening (ISB)
According to the Doppler equation, one should get the same velocity read-out at different Doppler angles
But, ISB artifact adversely affects this “law” of physics

## Slide 128: ISB is the result of wide Doppler apertures used in linear array transducers
Wide aperture linear array
Waveform broadened and peak frequency artificially elevated
Multiple elements on the face of the transducer are used for Doppler, the Doppler line represents the center elements.
Notes: Linear array transducers use numerous elements across the face of the transducer for Doppler purpose. This wide aperture effect causes a broadening of the spectral waveform and elevates the peak frequencies artificially. This effect is most pronounced at high Doppler angles, and is reduced at lower Doppler angles. The result can be an overestimation of measured peak velocity. This is a good reason for using Doppler Angles of 60 degrees or less.

## Slide 129: Because of the “wide” Doppler, the waveform is broadened and peak frequencies are artifically elevated
Wide aperture linear array
Waveform broadened and peak frequency artificially elevated
Notes: Linear array transducers use numerous elements across the face of the transducer for Doppler purpose. This wide aperture effect causes a broadening of the spectral waveform and elevates the peak frequencies artificially. This effect is most pronounced at high Doppler angles, and is reduced at lower Doppler angles. The result can be an overestimation of measured peak velocity. This is a good reason for using Doppler Angles of 60 degrees or less.

## Slide 130: Doppler Flow Phantom Phantom PSV 98 cm/s (red line)
Notes: The phantom motor is pulsating and timed to resemble blood flow. The red line is the peak velocity of the string phantom. 98 cm/sec. The Doppler spectrum above this line is due to ISB and is an artifact that causes the velocity readout to be 112 cm/sec.

## Slide 131: String Phantom Demo of ISB
35 degree angle to flow,
ISB spread is low.
70 degree angle to flow,
elevated spread and peak vel.
Peak vel = 59 cm/s
Avg vel = 45 cm/s
Peak vel = 78 cm/s
Avg vel = 46 cm/s
string
string
+
Notes: The frequency display of a moving string in the phantom should be thin, as the string is moving at a fixed, constant speed. The width, or spread is due to Doppler processing and intrinsic spectral broadening. When the Doppler angle to flow is changed to 70º the ISB is much greater, and peak velocity is also much higher. However, Note that the average velocity is similar! This is not affected by ISB.

## Slide 132: Intrinsic Spectral Broadening
Results in an overestimation of true blood flow velocity
ISB is minimal at low Doppler angles.
ISB is severe at angles ≥ 70º

## Slide 133: Range Gate Ambiguity- in High PRF
Manufacturers can double the PRF by transmitting before echoes have returned
This creates multiple, or phantom sample gates.
No problem, as long as the phantom SV don’t lie in an artery or vein.
Notes: When sample volume is placed deep, PRF is intrinsically lower, as the system must wait until the the echo returns before firing again. This can result in aliasing, and the PRF cannot be increased. Some manufacturers will allow a second pulse transmission before the original one arrives. The pulse internal is cut in half, and PRF can be doubled. When this occurs, a “range ambiguity” results. Let’s say the main SV is set at 4 cm, at 1540 M/sec it will take the 27 microsecond to reach the SV depth and 27 microseconds to return. So the pulse interval is 54 microseconds. In high PRF, a pulse is transmitted before the original pulse returns. The returning signal is being received in half the time, and twice the normal interval. Phantom gates will be displayed.
As long as the “phantom” gate is not in vessel, High PRF is effectiive.

## Slide 134: Multiple Range Gates in High PRF Mode